Published Date:	December 7 2023

Title:			GSI AudioStar Pro Version 2.1.0

Product(s):		GSI AudioStar Pro V2

Author:		GSI Product Management

# Introduction

The purpose of this bulletin is to announce an update to the GSI AudioStar Pro™ V2 firmware. There are also changes to the operating system and Configuration Application. These updates are in‐line changes and do not require a mandatory update to systems in the field; however, if a service person is on-site for any reason, GSI recommends that software is updated to the current release versions. Any new installation of a system should also be updated to the latest version. A copy of the updated software may be obtained from the GSI Extranet. This update includes the firmware to fix problems reported from the field and significant enhancements.

| Software Updates Description | Version |
| --- | --- |
| AudioStar Pro V2 Firmware | 2.1.0 (Build 498) |
| AudioStar Pro Config App | 2.1.0 (Build 63) |
| GSI Instrument Services | 2.12.0 (Build 35) |

# Enhancements

- Audible Contrast Threshold (ACT) test support (see below and attached materials)

- Added audiogram Vibrotactile symbol

- Weber Test support

- Configurable weighted PTA in Config App

- Device firmware can be updated from PC (future firmware)

- Option for color choice for Aided (HA and CI symbols)

- Ability to initiate patient response and VRA buttons from the QWERTY keyboard

- Can set the numbers of words to score for speech tests

- Added the Czech language as an option

# Bug Fixes

- Fixed a bug with the color of the masking table to reflect the ear of the stimulus

- Fixed Some dialogs that have text for the lower case that were cutting off the lower portions of the letters

- Fixed a Navigation Menu bug where the menu does not return to Signal (if Tone is default) or Test Type (if  Speech is default). Nav Menu remained on the last selected menu item

- Fixed bug where key navigation for select patient dialog when in modes other than patient not working

- Fixed a bug where the Hong Kong L/R sound field symbol was incorrect

- Fixed a bug where the patient response state changed and not registered while More menu was active

- Fixed issue here a blank wordlist favorite doesn't make a blank display screen

- Fixed bug where the Date Time dialog edit fields allowed to many characters, resulting in an overflow exception

- Fixed a bug where the Examiner button LED wasn't active when in Examiner mode

- Fixed issue of Talk Forward not working when in Patient Screen and Examiner dialog is open

- Fixed a bug with the audiogram refresh for connecting line display

- Fixed an out of memory exception when scrolling a list box

- Fixed an issue where the high-frequency headphone did not allow TENS testing

# ACT Test

This release of firmware provides support for the ACT test. The ACT diagnostic test is an above-threshold test that quantifies an individual’s real-world ability to hear in noise, with the intention of determining the level of help in noise needed. The ACT test results have good correlations with aided speech-in-noise performance. The ACT test is supported on AudioStar Pro V2 devices with the release of the version 2.1 firmware. The ACT test requires a license on the audiometer. New AudioStar Pro devices will be shipped with the ACT test license. For existing systems in the field, an upgrade to the firmware on the device and an update to the device license is required. Use the part number below to order an update. When ordering an upgrade option for the GSI AudioStar Pro, please provide the Serial Number and the License Number.

| Part Number | Description | Retail Price |
| --- | --- | --- |
| 8539785 | ACT License upgrade | $500.00 |

# Update Procedure

It is important to note that previously cached test results may not be usable after the update. Therefore, transfer all test results to the PC (i.e., GSI Suite, AudBase, etc.) prior to performing the update. Then go to the Patient screen, select Delete Sessions, and Delete All stored sessions.

NOTE: During the update, do not power down or use the instrument. If you power down the

instrument during the update, it will corrupt the device requiring it to be sent back to the

factory for repair.

## Update Firmware and OS:

- Save the firmware file for the audiometer onto the root of a USB drive.

- Insert USB drive into one of the USB ports on the audiometer.

- Press ConfigureUpdateDevice. The audiometer will re‐start after the update is complete.

NOTE: If the audiometer does not re-start and an error message appears, re-start audiometer manually.

- Confirm the firmware has been updated via the Configure screen (App version 2.1.0 Build 498). The OS version should be 2031, 2.2, 2031.

# Questions

For further assistance, please contact Grason-Stadler at (952) 278-4402, (800) 700-2282, or info@grason-stadler.com.

Thank you for your continued support of GSI.

800A-170

800A-121